## Slide 1: Changes in ISO/IEC 17043:2023 S. Subramanian

## Slide 2: Main changes are
harmonization with ISO/IEC 17025:2017, including technical requirements and structure;
harmonization with ISO 13528:2022 in terms of terminology;
incorporation of requirements from ISO/CASCO PROC 33;
inclusion of the requirement that testing activities, calibration activities and proficiency testing item production conform to the relevant requirements of appropriate ISO conformity assessment standards;
deletion of Annex C and revision of Annexes A & B.

## Slide 3: Clause 4.1 Impartiality – Similar to ISO/IEC 17025:2017
New standard has sub clauses from 4.1.1 to 4.1.6
Practically identical to clause 4.1 of ISO/IEC 17025:2017 (although 17025 standard has 5 sub clauses 4.1.1 to 4.1.5). No changes in the contents

## Slide 4: Clause 4.2 Confidentiality – Similar to ISO/IEC 17025:2017
New standard has sub clauses from 4.2.1 to 4.2.5 (one additional clause)
No changes in the contents of 4.2.1 to 4.2.4
But 4.2.5 is an additional requirement on confidentiality of identity of participant – Shifted from Clause 4.10.1 of old 17043 std.
Deleted clauses 4.10.3 & 4.10.4 of old 17043 std.

## Slide 5: Clause 5 Structural requirements – Similar to ISO/IEC 17025:2017
New standard has sub clauses from 5.1 to 5.7
No changes in the contents

## Slide 6: Clause 6.1 General– Similar to ISO/IEC 17025:2017
New standard has sub clauses from 6.1.1 to 6.1.3 (two additional clauses)
Addition-1: 6.1.2 Measurements or tests conducted under the responsibility of the PT provider, related to PT item characterization or for assessing homogeneity and stability, shall be conducted in accordance with the relevant requirements of ISO/IEC 17025 (excluding Management System requirements)
Addition-2: 6.1.3 Where the PT item is a material that meets the definition of “reference material”, it shall be produced under conditions that meet the relevant requirements of ISO 17034 (excluding Management System requirements)

## Slide 7: Clause 6.2 Personnel – Some changes in old 17043 standard - Clause 4.2
New standard has sub clauses from 6.2.1 to 6.2.7
Employment of temporary personnel given in 4.2.3 of old 17043 standard is not mentioned. No mention about Advisory Group explicitly in this clause.
Authorization of personnel given in 4.2.4 of old 17043 standard is changed. Authorization is specified only for 5 activities against 11 activities given in old 17043 standard
Training requirements given in the old 17043 std. in clauses 4.2.6 & 4.2.7 are deleted in the new std. Training requirement is given briefly.
New standard looks more like ISO/IEC 17025:2017 – Clause 6.2

## Slide 8: Clause 6.3 Facilities & environmental conditions – Some changes in old 17043 standard - Clause 4.3
New standard has sub clauses from 6.3.1 to 6.3.5
Performance characteristics of lab method and equipment given in 4.3.6 of old 17043 standard is deleted.
Slight re arrangement of requirements given in old 17043 standard

## Slide 9: Clause 6.4 Externally provided products and services – Similar to ISO/IEC 17025:2017
New standard has sub clauses from 6.4.1 to 6.4.6
Old 17043 standard two clauses 5.5 – Subcontracting services & 5.6 Purchasing services and supplies are combined
Looks like clause 6.6 of ISO/IEC 17025:2017, but some additional requirements are given by shifting the clauses 5.5.1, 5.5.2, 5.5.3 & 5.5.4 from the old 17043 standard as clauses 6.4.1, 6.4.2, 6.4.3 and 6.4.6 in the new 17043 standard

## Slide 10: Clause 7.1.1 Review of requests, tenders & contracts (Old standard clause 5.4)
New standard has sub clauses from 7.1.1.1 to 7.1.1.5
Addition: Review to ensure that the the objectives of the PT scheme are sufficiently defined and in agreement with the customers' needs;
Old 17043 standard clauses 5.4 – retained with slight re-arrangement

## Slide 11: Clause 7.1.2 PT Scheme communication (Old standard clause 4.9)
New standard has sub clauses from 7.1.2.1 to 7.1.2.3
Old 17043 standard clauses 4.9 – retained with 2 clauses shifted to other clauses in the new standard :
Old 17043 standard clause 4.9.3 regarding appeal by participants is shifted to clause 7.7 of the new standard (Handling of appeal) + Old 17043 standard clauses 4.9.5 regarding issue of statement on performance or participation of participants is shifted to clause 7.4.3 of the new standard (PT reports)
In the “protocol” to be sent to prospective participants added the following : (a) Objectives, (b) criteria for determining assigned value and the evaluation of performance, and (c) critical timeliness

## Slide 12: Clause 7.2.1 Design & Planning - General (Old standard clause 4.4.1)
New standard has sub clauses from 7.2.1.1 to 7.2.1.3
2 clauses of Old 17043 standard shifted and one clause deleted:
(a) 4.4.1.2 Planning not be sub contracted – Shifted to 6.4.1 in new standard
(b) 4.4.1.4 “Access” to technical expertise & experience – Shifted to 6.2.1 in new standard
(c) 4.4.1.5 Use of technical expertise (of advisory Group) – deleted
Addition 1: NOTE under 7.2.1.1 of new standard : When designing and planning the PT scheme, the relevant standards and requirements specific to the objectives of the PT scheme can be considered, e.g. ISO/IEC 17025, ISO 15189, ISO/IEC 17020. Safety and ethical issues can also be considered.

## Slide 13: Clause 7.2.1 Design & Planning - General (Old standard clause 4.4.1)
      Addition 2: 7.2.1.2 of new standard : When a PT provider intends to introduce significant changes to activities which can affect the validity of the PT scheme, the PT provider shall identify and manage the risk to ensure the validity of the PT scheme is maintained. Examples of significant changes are new approaches for PT item production, assessment of homogeneity and stability, determination of the assigned value, statistical analysis and new types of PT activities.
Addition 3 : in the design (plan) in new standard 7.2.1.3:
(q) the treatment of results from different measurement or test methods, where permitted by the PT scheme

## Slide 14: Clause 7.2.1 Design & Planning - General (Old standard clause 4.4.1)
Addition 4 : in the design (plan) in new standard 7.2.1.3:
NOTE -under (p) : Assigned values can have uncertainty contributions from sources in addition to the uncertainty of measurement results used for characterization, such as in-homogeneity and instability, and interlaboratory differences if more than one laboratory is used for characterization.
Minor changes in the design :
“Co-ordinator” replaced with “personnel”
“Precautions” to prevent collusion replaced with “arrangements”
“Selection of measurands … in PT round” replaced with “description of activities to be performed”
Actions to be taken in the case of lost or damaged PT items – added “delayed” also

## Slide 15: Clause 7.2.2 Statistical design (Old standard clause 4.4.4)
New standard has sub clauses from 7.2.2.1 to 7.2.2.3
Under 7.2.2.3 following 2 sub clauses added:
(f) the procedures to be used to treat participant results from different measurement or test methods which are not technically equivalent, where permitted by the PT scheme;
(g) whether the measurement uncertainty of participant results shall be reported and how it will be used to evaluate the participant's performance;

## Slide 16: Clause 7.2.3 Determination of assigned value (Old standard clause 4.4.5)
New standard has sub clauses from 7.2.3.1 to 7.2.3.5
Under 7.2.3.1 following change is made:
Where applicable, procedure (for determination of assigned value) shall take into account the metrological traceability and uncertainty required to demonstrate that the PT scheme is fit for its purpose. “Where applicable” is added now.

## Slide 17: Clause 7.3.1 Production of PT items (Old standard clause 4.4.2)
New standard has sub clauses from 7.3.1.1 to 7.3.1.3
Change :
Old 17043 standard clause 4.4.3 regarding use of PT items matching with items & materials encountered in routine testing is given as a note under clause 7.3.1.2 of the new 17043 standard
b) Minor additions in clause 7.3.1.3 regarding preparation and manipulation of PT items by participants - handling, storage, environmental conditions where applicable

## Slide 18: Clause 7.3.2 Homogeneity & stability assessment of PT items (Old standard clause 4.4.3)
New standard has sub clauses from 7.3.2.1 to 7.3.2.6
(i) Additions:
Note 2 under 7.3.2.3 : Different approaches for the assessment of homogeneity and stability, including situations where experimental study is not feasible, are described in Annex B of this document, in ISO 13528 and in ISO Guide 35.
New clause 7.3.2.4: Where experimental evidence is needed to assess homogeneity or stability of the PT item (or both), the PT provider shall use appropriate methods to assess the homogeneity and stability of the PT item.

## Slide 19: Clause 7.3.2 Homogeneity & stability assessment of PT items (Old standard clause 4.4.3)
(ii) Deletion:
Note 2 under old 17043 std. clause 4.4.3.3 regarding not carrying out homogeneity testing prior to distribution is deleted.
(iii) Shifting:
Old 17043 standard clauses 4.3.6 regarding circumstances where homogeneity & stability testing is not feasible is shifted as note 2 under the new standard clause 7.3.2.3

## Slide 20: Clause 7.3.3 Handling and storage of PT items (Old standard clause 4.6.2)
New standard has sub clauses from 7.3.3.1 to 7.3.3.4
Addition:
In the new std. clause 7.3.3.3 “prior to distribution” is added as shown below:
      7.3.3.3 The condition of stored PT items shall be assessed at specified intervals or prior to distribution in order to detect possible deterioration.

## Slide 21: Clause 7.3.4 Packaging, labelling & distribution of PT items (Old standard clause 4.6.3)
New standard has sub clauses from 7.3.4.1 to 7.3.4.5
Addition:
In the new std. clause 7.3.4.3 “to ensure the validity of the PT item” is added as shown below:
      7.3.4.3 In PT schemes where participants are required to transport the PT items to other participants, or return them to the PT provider, documented instructions for this transport, to ensure the validity of the PT item, shall be supplied.

## Slide 22: Clause 7.3.5 Instructions for participants (Old standard clause 4.6.1)
New standard has sub clauses from 7.3.5.1 to 7.3.5.2
Addition
New sub clause(g) under 7.3.5.2 as given below:
(g) specific instructions on providing details concerning the measurement or test method used by the participant, where a single specific measurement or test method is not required;
(ii) Added the text “that would not be considered ………. the PT scheme” in the new std. clause 7.3.5.2 (c) as given below:
(c) instructions for preparing or conditioning, or both, of the PT items before conducting the measurements or tests that would not be considered part of a laboratory’s usual expected practices, unless these activities are part of the PT scheme;

## Slide 23: Clause 7.4.1 Data analysis (Old standard clause 4.7.1)
New standard has sub clauses from 7.4.1.1 to 7.4.1.6
Addition
Clause 7.4.1.4 added in the new std. as given below:
7.4.1.4 The PT provider shall have procedures for treatment of results from different measurement or test methods, where the PT scheme allows participants to use different measurement or test methods.
Shifting
Old 17043 standard clauses 4.7.1 regarding “validation of software” is shifted to the new Std. Clause 7.5.2.2

## Slide 24: Clause 7.4.3 PT reports (Old standard clause 4.8)
New standard has sub clauses from 7.4.3.1 to 7.4.3.7
Addition : Clause 7.4.3.6 added in the new std. as given below:
      7.4.3.6 When issuing an amended report to a subset of participants, an analysis of the potential impact on the other participants for that PT scheme and/ or PT round shall be made to ensure there is no influence on the general performance of the other participants.
Shifting:
Old 17043 standard clauses 4.9.5 regarding “issue of statement of performance or participation to participants ” is shifted to the new Std. as Clause 7.4.3.7
Deletion:
Name & contact details of PT Coordinator deleted.

## Slide 25: Clause 7.5.1 Technical records (Old standard clause 5.13.2)
New standard has sub clauses from 7.5.1.1 to 7.5.1.3
New standard looks more like ISO/IEC 17025:2017 – Clause 7.5
Clause 7.5.1 of ISO/IEC 17025:2017 is divided in to 2 clauses – 7.5.1.1 & 7.5.1.2 in the new std., while clause 7.5.2 of ISO/IEC 17025:2017 is named as 7.5.1.3 in the new std.

## Slide 26: Clause 7.5.3 Surveillance of the processes NEW ADDITION
New standard 7.5.3 has no sub clause & same is given below:
7.5.3 The PT provider shall have a procedure to ensure the validity of the PT scheme. Surveillance activities shall be planned and reviewed [see also 8.9.2 item n)], and the resulting data shall be recorded for the continuous improvement process.
NOTE: Depending on the PT scheme, surveillance activities can include:
evaluation of externally provided products and services;
use of reference materials or other control items;
the transmission of results from participants;
control of statistical conditions to confirm the validity of performance evaluation;
checking of reports;
for continuous schemes, comparisons against previous PT rounds.

## Slide 27: Clause 7.6 Handling of complaints (Old standard clause 5.8)
New standard has sub clauses from 7.6.1 to 7.6.9
New standard looks more like ISO/IEC 17025:2017 – Clause 7.9
Following additions (blue text) to clause 7.9 of ISO/IEC 17025:2017 included in clauses 7.6.6 & 7.6.7 of new std. ISO/IEC 17043:2023:
    7.6.6 Investigation and resolution of complaints shall not result in any discriminatory actions.
7.6.7 The resolution of complaints shall be made by, or reviewed and approved by, persons not involved in the subject of the complaint in question. Where resources do not permit this, any alternative approach shall not compromise impartiality

## Slide 28: Clause 7.7 Handling of appeals (Old standard clause 5.8)
New standard has sub clauses from 7.7.1 to 7.7.7
New standard looks more like ISO/IEC 17025:2017 – Clause 7.9
The following addition is included as a note under clause 7.7.1 of the new std. ISO/IEC 17043:2023:
NOTE: PT providers that only have PT schemes using purely statistically derived evaluation procedures do not usually handle appeals. Appeals concerning performance evaluations can be addressed as a complaint.

## Slide 29: Clause 8.1 General requirements - Similar to ISO/IEC 17025:2017
New standard has sub clauses from 8.1.1 to 8.1.4
Options 1 and 2 of ISO/IEC 17025:2017 -Clause 8.1 are not given explicitly here
Addition:
Evidence of commitment of management (given in 8.2.3 of ISO/IEC 17025:2017) is shifted to this clause as 8.1.4 of ISO/IEC 17043:2023

## Slide 30: Clause 8.2 Management system documentation- Similar to ISO/IEC 17025:2017
New standard has sub clauses from 8.2.1 to 8.2.3
Establishing and documenting policies and objectives (given in 8.2.1 of ISO/IEC 17025:2017) is deleted
Evidence of commitment of management (given in 8.2.3 of ISO/IEC 17025:2017) is shifted to this clause as 8.1.4 of ISO/IEC 17043:2023